1. Devanning and placement

The VSM packing box is made of plywood，there may be unprocessed rags on the body of the plywood, so please wear gloves for protection during disassembly.

Check if the box is damaged, water seepage, liquid leakage, cracking, or screw loose before unpacking the box.

Upper cover plate will be opened firstly when unpacking, then remove the side panels. The disassembling sequence of the side panels depends on your habits.

Note: Do not destroy the box.

Note: Do not use brute force to knock, break or open the package during unpacking.

Caution: Items in the box may be scattered when the side panels are dismantled .

After opening all the boxes, please check whether the box and accessories are missing according to the packing list (the water cooler and dehumidifier can be carried by manpower, while the vibration head frame and electromagnet need to be carried by machine).

The whole machine is packed in four boxes separately. First box is filled with electromagnet, second box is filled with cabinet, third box is filled with accessories (including vibration head frame, etc.), fourth box is filled with water cooler.

Notice: There is a packing list on each box.

Please pay attention to third box. After opening the box, you should take out and check the package firstly, then check the accessories one by one according to the packing list.

Confirm the installation position, refer to the layout chart, take out the electromagnet, electromagnet should be carried by machine.

The parts in the box are precision parts, avoid vibration when landing, so as not to affect the use.(Note : the position of the electromagnet should be close to the power supply, the bearing capacity of the ground and the level of flatness refers to “condition of installment”)

Remove the ring of electromagnet and place the four feet margins roughly near the electromagnet.

The cabinet is placed on one side near the electromagnet, and the chiller is placed on the other side of the electromagnet.

(There is a brake at the bottom of the chiller, lift the brake firstly when moving, and then lay down the brake to prevent moving after the position is determined)