11. Preparation of powder sample before testing

The powder sample is usually wrapped with Teflon tape, which is PTFE film, for measurement. Cut about 5 cm length of Teflon type with scissors, and put it on electronic weighing scale; after the scale reading is stable, deduct the tare; put a proper amount of powder sample on one end of Teflon tape, of which the distance is about 1 cm to the end. Ensure to take the small quantity of powder sample, or if too much, the Teflon tape could not wrap the powder sample completely. Now the scale reading of the sample quality is 64.5 mg.

After writing down the reading, take out the Teflon tape with powder sample from the weighing scale, wrap the powder sample in Teflon tape by tweezers. During wrapping the sample by tweezers, try to make each layer of the tape to evenly wrap the previous layer, and prevent the powder sample from leakage；Then wrap the sample in Teflon tape by handers, try to stretch the tape out a little longer, and wrap the sample tightened, now the powder sample is ready for measurement.

The above preparation method of powder sample is only suitable for the sample, of which coercive force is not high. Due to the inaccuracy measuring result of coercivity is not accuracy, therefore