3. Position adjustment of electromagnet and vibrating head frame (including vibrating rod)

Place the gradienter on the bottom panel of the electromagnet;

Adjust the electromagnet’s leveling feet with an adjustable wrench;

Judge whether the electromagnet is in the horizontal according to the gradienter’s indication;

The bubble is in the center of the gradienter, which indicates that the electromagnet has been adjusted to be in the horizontal;

After the electromagnet in the horizontal, it can start to adjust the position of vibrating head frame;

Place the gradienter on the vibrating head frame;

Adjust the leveling feet of vibrating head frame, and make the bubble is in the center position of the gradienter;

Observe the position of the vibrating rod’s protective sleeve, and adjust the vibrating head frame in order to put the vibration rod’s protective sleeve roughly in the center of the magnetic field;

About the specific position, pls refer to the position diagram of electromagnet and vibrating head frame (Suggest to pause the video and see the diagram);

After adjusting the horizontal position of the vibrating head frame, re-adjust each of the leveling feet of the vibrating head frame separately and make the bubble in the center position of the gradienter;

After adjusting the position of the vibrating head frame, then it can start to install the vibrating rod;

Following the operation shown in the previous video, unscrew the locking cap of the vibrating rod;

Insert the vibrating rod in the shown position;

Tighten the locking cap of the vibrating rod;

Install the dust cover;

Now the vibration system assembly is finished.

Summary

1. Firstly, adjust the electromagnet to be in the horizontal;

2. Then, adjust the vibrating head frame to be in the horizontal;

3. Next, re-adjust each of the levelling feet separately and make the vibrating head frame to be in the horizontal;

4. Then, make sure the vibrating head frame is in the horizontal again;

5. Finally, install the vibrating rod.