5. Wiring Connection

1.Take out the COM/USB cable, and insert the nine-pin terminal into the connection port of the PC’s cabinet, and then tighten the screws of the both sides on the nine-pin terminal;

Insert the other end of the COM/USB cable through the cable entrance hole on the top of the cabinet, find the matched COM/USB port inside the cabinet and tighten the cable clockwise.

2. Take out the OSC•OUT cable;

Insert the black end of the OSC•OUT cable into the cabinet through the cable entrance hole at the bottom of the cabinet;

Then, plug it into the OSC•OUT connection port on the cabinet;

The other end of the OSC•OUT cable is plugged into the matched OSC•OUT port on the vibrating head;

3. Take out the M_INPUT cable;

Insert the three-core end into the cabinet through the cable entrance hole at the bottom of the cabinet;

Plug the cable into the matched M_INPUT connection port on the cabinet;

The other end is plugged into the matched M_INPUT connection port of the electromagnet;

4. Take out the H_INPUT cable;

One end of the cable is a Hall Probe, which is wrapped in a plastic tube;

Remove the adhesive tape protective layer from the Probe;

Insert the connecting end of the cable into the cabinet through the cable entrance hole at the bottom of the cabinet;

Plug it into the matched H_INPUT connection port of the cabinet;

The other end, which is the Hall Probe, will be extended from the protective sleeve;

Insert the probe into the matched through hole of the probe holder;

Adjust the position of the probe, and make the tip of the probe is in the center of the magnetic field and make the probe near the pole face on one side;

Both sides of the probe holder are the rotatable supports, which can be used for adjusting the overall position of the probe;

5. Take out the electromagnet’s earth wire as shown in the video;

Both ends of the earth wire are the same cable lugs. Either end can be inserted into the cabinet through the cable entrance hole at the bottom of the cabinet;

Connect the earth wire to the grounding connection port in the cabinet with a hex wrench;

On the other end of the earth wire is linked to the electromagnet’s grounding connection port with a hex wrench;

6. Take out power supply cable of electromagnet as shown in the video;

Connect “+”positive terminal to “+” positive terminal connector on the cabinet, and tighten the screw on the connector with a slotted screwdriver;

Connect “-” negative terminal to “-” negative terminal connector on the cabinet, and tighten the screw on the connector with as lotted screw driver;

The other end of the power supply cable is linked to the electromagnet;

Unscrew the binding post, which the copper cable lug will be connected to, on the electromagnet;

Connect the copper cable lug marked with "+"the positive to "+" positive binding post, and screw the post tightly;

Link the copper cable lug marked with "-" the negative to "-" negative binding post, and screw the post tightly;

7. Take out the cabinet’s three-phase cable as shown in the video；

Both ends of the three-phase cable are the terminals. Plug either end of the cable into the cabinet through the cable entrance hole at the bottom of the cabinet;

“N” terminal is connected to the matched “N” terminal connector on the cabinet;

“PE” terminal is linked to the grounding connector on the cabinet;

“C” terminal is connected to the matched “C” terminal connector on the cabinet;

“B” terminal is linked to the matched “B” terminal connector on the cabinet;

“A” terminal is connected to the matched “A” terminal connector on the cabinet;

The wiring connection of the cabinet are finished.

Three terminals of "A", "B","C" will be plugged into the miniature circuit break(MCB) or air circuit break(ACB) in a positive phase sequence according to the actual situation on site.

Ensure "A" terminal of three phase cable, which has been plugged into the cabinet, is connected to the matched “A” terminal connector on MCB / ACB;

Ensure "B" terminal of three phase cable, which has been plugged into the cabinet, is connected to the matched “B” terminal connector on MCB / ACB;

Ensure "C" terminal of three phase cable, which has been plugged into the cabinet, is connected to the matched “C” terminal connector on MCB / ACB;

“N” terminal is connected to null line port of MCB / ACB;

“PE” terminal is connected to earth line port of MCB / ACB, and ensure the wiring connection firmly.

There are water inlet and water outlet on the back of chiller;

Unscrew the plugs of water inlet and outlet with an adjustable wrench;

8. Take out two taps of chiller.

Wrap PTFE around the threads on connectors for linking between the taps and the water inlet and outlet of chiller;

Then, screw a tap into water outlet;

Tighten the tap with an adjustable wrench;

Do the same operation for another tap;

Take out four plastic washers from the zip lock bag;

Place one washer at the connector of one water pipe;

Connect the water pipe to a tap;

Tighten the nut at the connector with an adjustable wrench;

Do the same operation for the other water pipe;

Place one of the rest washers at the connector at the other end of water pipe;

Connect the above water pipe to a water connector on the electromagnet, and there is no difference between water inlet and water outlet of electromagnet;

Tighten the nut at the water connector with an adjustable wrench;

Do the same operation for the connector of the other water pipe.

Summary

1. COM/USB cable connection;

2. OSC•OUT cable connection;

3. M_INPUT cable connection;

4. H_INPUT cable connection;

5.  Electromagnet’s earth wire connection;

6. Electromagnet’s power supply cable connection;

7. Cabinet’s three-phase cable connection;

8. Connect water pipes between chiller and electromagnet.