7. Sample Test Operating Procedures

After calibration, unload the standard nickel ball sample;

Load the tested sample to the vibrating rod, and prepare to test the sample;

Before testing sample, click “Display Off” item, and switch it to “Display On” item, then check and ensure the value of magnetic moment approaching 0 (M value should be less than 0.0005);

If not, the value of magnetic moment is reset to be zero again;

Click “User Name”, and create a folder under the directory of Data file, under D disk, then input a file name into the right text item of “Sample No.”;

This example will use a ball-shaped permanent magnetic sample, and put the sample into the standard sample chamber, next put a small ball of absorbent cotton, which plays a role of buffer, into the sample chamber too, then load the sample chamber to the vibrating rod; As this sample is the ball-shaped sample, of which the position is as same as the standard Nickel ball sample, therefore, it is not necessary to adjust the Nek Area, and just test the sample;

In this example, choose "Test normal m (Hext) loop" item as a measurement mode to test Hysteresis Loop Curve;

Set the parameters of the tested sample according to the sample’s actual characteristics;

Before testing sample, do the following operating steps again: click “Display Off” item, and switch it to “Display On” item, then double check if the value of magnetic moment is approaching 0 (M value should be less than 0.0005);

If not, the value of magnetic moment is reset to be zero again;

Click “Test Sample” item, and begin to test;

Testing;