9. Introduction of Original Data

Find the saved position of the measured files in PC and open it;

“*.13D” is the format of the measured files，and “Notepad” is the default open mode of the files;

The first and second lines are the description data;

The first number in the first line is mass value of the measured sample, and the unit is “mg”;

The second number in the first line is density value of the measured sample, and the unit is “g/cm^3”;

The third number in the first line is “Measurement Function”, which start from “0”, and the current test mode shows “2”;

That means the present measurement mode is “Test magnetization curve”;

In the drop-down menu of “Measurement Function”, “0” refers to “Fast test m (Hext) loop”, “1” refers to “Test normal m(Hext) loop”, “2” refers to “Test magnetization curve”, and so on;

The fourth number in the first line is demagnetization factor, if the mass and density have been input,

Then, the demagnetization factor will join in the calculation, the vale range is from “0” to “1”;

The first number in the second line is the temperature when the sample is measured, the unit is “K”; (Whether is under High-Temperature Mode, or Low-Temperature Mode);

The two numbers in the middle of the second line are both “0”, which are the meaningless values for room temperature VSM equipment;

The last data in the second line is “emu”, which is the unit of magnetic moment;

From the third line to the last line, there are two columns, which are the original data;

The left column is the value of magnetic field, and the unit is “Oe”;

The right column is the value of magnetic moment, and the unit is “emu”;

After the original data have been saved in this VSM software, these data could not be changed or modified;

There are only the values of mass, density, demagnetization factors can be changed and modified in this VSM software;

When the measured data need to be copied to EXCEL or other mathematical processing software;

It is ok to choose and separate them with commas.